## Summary for process checking(2)
Step | Process Module | Equipment name | Model/Size in URS | Qty in URS | URS | Note | Question1 | Answers1 | Respostas | Question2 | Answers 2
| General info |  |  |  | URS2, 5.1.74 | URS indicates: The mRNA production platform must be capable of producing at least 4 (four) grams of purified RNA per batch. | 1.Since different companies have different mRNA production processes, please describe in detail your specific production steps. If you have the process flow diagram (PFD), please share it with us. 2.Based on the requirements in the URS2, 5.1.74: Please confirm this means each IVT reaction batch produces 4g of purified mRNA? For annual production 750g of mRNA (15M doses and 50ug per does), it will need to produce 188 batches per year. This batch quantity is too large, so it is recommended to increase the batch production capacity in IVT reaction. | 1. pDNA Production (50 L Fermenter): Typical yield is 2–3 g of circular pDNA per batch, with each batch taking 3–4 days. Over 300 days, approximately 75 fermentation cycles can be completed, resulting in ~187 g of circular pDNA. After linearization (85% yield), this gives ~159 g of linear pDNA. 2. mRNA Production via IVT (50 L Reactor): Each IVT cycle takes 2 days and can process ~1.06 g of linear pDNA, yielding ~21.2 g of mRNA per cycle. Over 300 days, ~150 cycles can be run, producing a total of ~3,180 g (3.18 kg) of purified mRNA. | 1. Produção de pDNA (Fermentador de 50 L): O rendimento típico é de 2 a 3 g de pDNA circular por lote, com cada lote levando de 3 a 4 dias. Ao longo de 300 dias, aproximadamente 75 ciclos de fermentação podem ser concluídos, resultando em ~187 g de pDNA circular. Após a linearização (rendimento de 85%), obtém-se ~159 g de pDNA linear. 2. Produção de mRNA via IVT (Reator de 50 L): Cada ciclo de IVT leva 2 dias e pode processar ~1,06 g de pDNA linear, produzindo ~21,2 g de mRNA por ciclo. Ao longo de 300 dias, podem ser realizados ~150 ciclos, produzindo um total de ~3.180 g (3,18 kg) de mRNA purificado. | Based on the normal practise, working volume of SUT fermenter is 30L. Then the assumption is Circular pDNA: "0.6 g/L of crude pDNA can be achieved"->30L*0.6g/L=18g per batch "40% recovery rate"-> 18g*0.4=7.2g per batch "After linearization (85% yield)"-> 7.2g*0.85=6.12g per batch; "approximately 75 fermentation cycles can be completed" -> 6.12g*75 batch=459g per year “1.06 g of linear pDNA, yielding ~21.2 g of mRNA per cycle”-> Plasmid convert to mRNA=1:20; So mRNA per year 459g*20=9180g, After purification 60% recovery rate -> 9810g*0.6=5508g per year (target 750g mRNA per year) | ok
pDNA | Media preparation | Media preparation system |  |  | URS1, Page 5 and 5.1.217-5.1.264 | No volume of Single Use Mixer indicated | 1.What is the volume of the preparation system? 2.How many units are needed for each volume of the preparation system? | 10-100L | 10-100L | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
| Fermentation | Fermentation system | No volume of bioreactor indicated | 2 | URS1, Page 4 and 5.1.265-5.1.364 | Bioreactors for Bacterial Growth and pDNA Linearization, No volume of bioreactor indicated | 1.According to the public data, the expression level of the plasmid through fermentation is 0.5-1 g/L, and the recovery rate of plasmid purification is about 40%. Based on URS, for each batch of production of 0.5 g pDNA, 4-8 L of fermentation broth is required. Please confirm whether this production capacity is suitable for your production. 2.What is the concentration of the expression plasmid for E. coli fermentation in g/L? 3.What is the volume of the fermentation system? 4.Does the Qty 2 refer to two different-sized fermentation systems or two of the largest production-level fermentation systems? | In a 50 L fermenter, dry biomass reaches approximately 75 g, yielding an average of 2.5 g of circular pDNA per batch. With a 40% recovery rate, each batch produces about 1.0 g of purified pDNA, allowing for 75 cycles per year and a total of 75 g of purified circular pDNA. After linearization (85% yield), this results in ~63.75 g of linear pDNA. Using a 50 L IVT reactor, each cycle (2 days) processes ~0.425 g of linear pDNA, producing ~8.5 g of mRNA per cycle, totaling ~1.28 kg of mRNA over 150 cycles annually. With optimized cultivation, up to 0.6 g/L of crude pDNA can be achieved, equating to 30 g per batch. The 50 L scale is deemed sufficient, and ideally, the single-use bioreactor should support flexible operation with 10 L and 50 L disposable bags.
| Buffer preparation | Buffer preparation system |  |  | URS1, Page 5 and 5.1.217-5.1.264 | No volume of Single Use Mixer indicated | What are the Volume and quantity? | 10-100L | 10-100L | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
| Collection | Alkaline lysis system |  |  | URS1, (5.1.33.) | an alkaline lysis system No content of parameter requirement indicated in URS | 1.What is the volume of the material to be lysed? 2.What is the solid content? | 1.5 g/L in dry biomass extract | 1,5 g/L em extrato de biomassa seca
|  | Mixer |  |  | URS1, Page 6 | No content of parameter requirement indicated in URS | After each step of clarification and purification process, the material needs to be temporarily stored in a mixer. Please confirm the quantity and volume. | volume between 10 and 50L | volume entre 10 e 50L | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | Continuous flow centrifuge |  |  | URS1, Page 5 | To separate cellular debris from the supernatant containing plasmid DNA after lysis, if applicable. No content of parameter requirement indicated in URS | 1.What is the volume of the material? 2.What is the solid content? 3.Base on our experience, TFF system is capable to perform this process. Is TFF acceptable? | 1. 10-50L 2. <1.5g/L 3. TFF may be acceptable for both microfiltration, diafiltration, and concentration. | 1. 10-50L 2. <1,5g/L 3. O TFF pode ser aceitável para microfiltração, diafiltração e concentração. | Customer agreed to use TFF system in this step.
|  | Storage system |  |  | URS1, Page 6 | No content of parameter requirement indicated in URS | What are the Volume and quantity? | Between 10 and 50L | Entre 10 e 50L | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
|  | Ultrafiltration system | Flow rate: 10 to 800 L/h | one or more | URS1, Page 5, and 5.1.97-5.1.216 | No filtration area indicated | Both before and after chromatography, the UF system needs to be used. It is optional to use the hollow fiber TFF system for harvesting the bacteria. Please confirm how many sets of UF systems are needed in your process? How much membrane area is required for each step of UF? What is the final volume after concentration? | Yes, but we can use the same system for different stages. The membrane area can vary, but for these volumes, 2 m² can cover all stages, based on calculations made by a specific supplier. Ideally, the supplier should demonstrate this rationale to us with the specifications of their membranes. | Sim, mas podemos usar o mesmo sistema para diferentes etapas. A área da membrana pode variar, mas para esses volumes, 2 m² podem cobrir todas as etapas, com base em cálculos feitos por um fornecedor específico. O ideal é que o fornecedor nos demonstre essa justificativa com as especificações de suas membranas. | Customer accepts the equipment is used for hollow fiber, in the process of harvesting, ultrafiltration after lysis, and ultrafiltration after chromatography
|  | Chromatographic system | chrom system: 15L/H-600L/H |  | URS1, Page 6 and 5.1.365-5.1.442 | Remove RNA, oDNA and endotoxin. It is usually a 3-step chromatography process. The sizes of the three chromatography columns will be different. Based on our experience, the volume of the first step SEC chromatography packing resinl is calculated to be 20L (the column height is greater than 50cm), the volume of the second step chromatography packing resin is 3L, and the volume of the third step chromatography packing resin is 2.5L. This result is quite different from the URS you provided (volume 36L, column height 28.5cm). We propose 1. (Diameter 200mm; Bed volume 15.7-22L), 2. (Diameter 140mm; Bed volume 0.8-4.6L), 3. (Diameter 140mm; Bed volume 0.8-4.6L) | Please confirm whether 36L of resin is required for the production of 0.5g of pDNA. The resin loading capacity is less than 0.014g/L, seems too low. | According to our tech transfer partners, a single ion exchange step would be sufficient. We've left two options, including the hydrophobic interaction step, for future development if necessary. The supplier can suggest other scales and processes and explain the rationale, which we'll evaluate. | De acordo com nossos parceiros de transferência de tecnologia, uma única etapa de troca iônica seria suficiente. Deixamos duas opções, incluindo a etapa de interação hidrofóbica, para desenvolvimento futuro, se necessário. O fornecedor pode sugerir outras escalas e processos e explicar a justificativa, que avaliaremos. | Customer confirm the flow patch of all chrom systems shall be SUT
|  | Column | Column: 400mm dia | 2 or more | URS1, Page 6 and 5.1.365-5.1.442 |  |  |  |  | Normally the resin will be used multiply. Customer accepts the column to br used multiply based on CIP. Customer also expects a quotation of resin price to evalute the operation cost if using the resn only one time for each batch.
|  | Sterilizing filtration system |  |  | URS1, Page 6 | No content of parameter requirement indicated in URS | 1.The volume of the material to be processed? 2.Filtration flowrate？ | The volume will be between 1 and 50 L and the filtration rate can be calculated based on this value. | O volume estará entre 1 e 50 L e a taxa de filtragem pode ser calculada com base nesse valor.
|  | Storage system |  |  | URS1, Page 6 | No content of parameter requirement indicated in URS | For Buffer storage. What are the Volume and quantity? | Between 0.5 and 50L | Entre 0,5 e 50L | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
mRNA | Buffer preparation | Buffer preparation system |  |  |  | Buffer preparation | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor 6.output process：☐0.45μm+0.22μm filter ☐0.22μm+0.22μm filter ☐0.45μm+0.22μm+0.22μm filter ☐No need | It could be a single-use buffer preparation system that serves the entire plant. Temperature control should be performed by a TCU. Weight measurement should ideally be done by a load cell; the supplier can propose other solutions. Filtration could be 0.45µm in different processes, with a final filtration ideally of 0.22µm. | Poderia ser um sistema de preparação de tampão de uso único que atendesse toda a planta. O controle de temperatura deveria ser realizado por uma TCU. A pesagem deveria ser feita, idealmente, por uma célula de carga; o fornecedor pode propor outras soluções. A filtração poderia ser de 0,45 µm em diferentes processos, com uma filtração final ideal de 0,22 µm. | Customer indicated the linearization volume will be 1/10 volume of the fermentation, e.g. 50L Fermentation will have 5L Linearization
| Digestion | Wave |  |  |  | Enzymatic digestion | What is the volume? | Between 500 mL and 10 L | Entre 500 mL e 10 L
|  | Mixer |  |  |  | Intermediate material solution storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | Between 500 mL and 10 L, or as required for the next in vitro transcription step | Entre 500 mL e 10 L, ou conforme necessário para a próxima etapa de transcrição in vitro | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | Chromatographic system |  |  |  | Remove non-linearized plasmids and endonucleases | Equipment quantity ? ； Please check below requirement of the system: ☐Single - use chromatography system，☐Stainless steel chromatography system. ☐ Gradient function； ☐ Chromatography system: flow rate： L/h； Inlet quantity： Pcs， Outlet quantity： Pcs； ☐ Pre-filter； Detectors： ☐ Bubble sensor at the sample inlet；☐ Bubble sensor at the pre-column； ☐ UV sensor： pcs； Brand ; ☐ single wavelength, nm;☐ dual wavelength, nm; ☐ full wavelength, nm; ☐ Pressure sensor: pcs, location: ☐ post pump, ☐ pre-column, ☐ post-column; ☐ pH sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Conductivity sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Temperature sensor pcs, location . ☐ Explosion - proof function; The salt concentration of buffer and sample in each step mol，pH ； Other requirement: . | Quantity: 2 dedicated chromatographic systems with capacity for continuous or campaign operation, depending on production volume. (ion exchange and hydrophobic iteration) System Type: Single-use system System Flow Rate: 5–20 L/h Inlets and Outlets: • Inlets: 2–4 ports (sample, buffers) • Outlets: 2 ports (collected fractions and waste) Pre-filter: Yes, to protect the column from particles Sensors and Detectors • Bubble Sensor: At the sample inlet and before the column • UV Sensor: 1 unit o Type: Dual wavelength (e.g., 260/280 nm and/or fluorescence) • Pressure Sensor: 3 units o Location: Post-pump, Pre-column, Post-column • pH Sensor: 2 units o Location: Pre-column, Post-column • Conductivity Sensor: 2 units o Location: Pre-column, Post-column • Temperature Sensor: 1 Unit o Location: On the outlet line or on the column • Explosion-proof function: Yes • Buffer salt concentration and pH: o Salt: 0.1–1.0 mol/L NaCl or K Acetate o pH: 4.0–10.0, depending on the step (equilibrium, wash, elution) Operation type • Automatic Column dimensions: Internal diameter: According to the supplier's proposal, maintaining final productivity Height: According to the supplier's proposal, maintaining final productivity Chromatographic resin Particle size: 0–100 μm Packaging equipment - The column and its purpose must be dedicated to each type of plasmid or genetic material. | Quantidade: 2 sistemas cromatográficos dedicados com capacidade de operação contínua ou em campanha, dependendo do volume de produção. (troca iônica e iteração hidrofóbica) Tipo de Sistema: Sistema de uso único Vazão do Sistema: 5–20 L/h Entradas e Saídas: • Entradas: 2–4 portas (amostra, tampões) • Saídas: 2 portas (frações coletadas e resíduos) Pré-filtro: Sim, para proteger a coluna de partículas Sensores e Detectores • Sensor de Bolhas: Na entrada da amostra e antes da coluna • Sensor UV: 1 unidade o Tipo: Comprimento de onda duplo (ex.: 260/280 nm e/ou fluorescência) • Sensor de Pressão: 3 unidades o Localização: Pós-bomba, Pré-coluna, Pós-coluna • Sensor de pH: 2 unidades o Localização: Pré-coluna, Pós-coluna • Sensor de Condutividade: 2 unidades o Localização: Pré-coluna, Pós-coluna • Sensor de Temperatura: 1 unidade o Localização: Na linha de saída ou na coluna • Função à prova de explosão: Sim • Concentração de sal e pH do tampão: o Sal: 0,1–1,0 mol/L NaCl ou Acetato de Potássio o pH: 4,0–10,0, dependendo da etapa (equilíbrio, lavagem, eluição) Tipo de operação • Automático Dimensões da coluna: Diâmetro interno: Conforme proposta do fornecedor, mantendo a produtividade final Altura: Conforme proposta do fornecedor, mantendo a produtividade final Resina cromatográfica Tamanho de partícula: 0–100 μm Equipamento de embalagem - A coluna e sua finalidade devem ser dedicadas a cada tipo de plasmídeo ou material genético. | Normally no explosion proof needed in this process as no ethanol is used in this step of mRNA purification. Customer will confirm whether to have explosion proof in this chromtographic system | Not required
|  | Column |  |  |  |  | Please check below requirement of the system: ☐Automatic；☐Manual； Column tube diameter mm； Column tube height mm； Column tube material: ☐Acrylic; ☐Stainless steel; Particle size of the chromatography resin μm； Mesh size: ☐10μm; ☐20μm; ☐Others: μm; Maximum operation pressure bar； Packing equipment： ☐ Automatic packing,☐ Manual packing； ☐ Automatic packing station; ☐ Slurry tank；volume: L; |  |  | Normally the resin will be used multiply. Customer accepts the column to br used multiply based on CIP. Customer also expects a quotation of resin price to evalute the operation cost if using the resn only one time for each batch.
|  | UF system |  |  |  | Concentration and diafiltration | Please check below requirement of the system: ☐Semi-automatic; ☐Automatic； Initial volume： L；Final volume L； Max. flow rate L/h； ☐Cassette，brand ，model ； ☐Hollow fiber，brand ，model ； Max membrane area： m2; System: ☐ Integrity test port reserved; ☐ Online integrity test module； ☐ Pressure sensor； ☐ Temperature sensor pcs，location: ☐ feed, ☐ retentate, ☐ permeate; ☐ pH sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Conductivity sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ UV sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Electromagnetic flowmeter, location: ☐ retentate, ☐ permeate; ☐ Mass flowmeter, location: ☐ retentate, ☐ permeate; ☐ Pipeline sampling valve, location: ☐ retentate, ☐ permeate; The salt concentration of the sample mol，pH ； ☐Organic reagents, concentration ； ☐Explosion - proof function； ☐TFF system circulating tank: Tank volume L，☐Weighing，☐Stirring，☐Tank thermal insulation layer；☐Temperature sensor; ☐Conductivity sensor; ☐ Sampling valve; ☐ Replenishment module: ☐ Replenishment inlet quantity ; Replenishment flow rate L/h； | The ultrafiltration (UF) system for concentration and diafiltration or microfiltration should preferably be automatic, with adjustable initial and final volumes depending on the process, and a maximum flow rate compatible with the production scale. It can use cassettes or hollow fibers, with an adequate membrane area (up to several m²), and include pressure, temperature, pH, conductivity, and UV sensors, as well as electromagnetic or mass flow meters, sampling valves, and an integrity testing module. The system can operate with organic reagents, buffers with defined saline concentrations and pH, and include a recirculating tank with agitation, thermal insulation, sensors, and a replenishment module with flow control. | O sistema de ultrafiltração (UF) para concentração e diafiltração ou microfiltração deve ser preferencialmente automático, com volumes inicial e final ajustáveis ​​dependendo do processo e vazão máxima compatível com a escala de produção. Pode utilizar cassetes ou fibras ocas, com área de membrana adequada (até vários m²), e incluir sensores de pressão, temperatura, pH, condutividade e UV, além de medidores de vazão eletromagnéticos ou mássicos, válvulas de amostragem e um módulo de teste de integridade. O sistema pode operar com reagentes orgânicos, tampões com concentrações salinas e pH definidos, e incluir um tanque de recirculação com agitação, isolamento térmico, sensores e um módulo de reabastecimento com controle de vazão.
|  | Storage system |  |  |  | Buffer storage | 1. What are the Volume and quantity Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | The buffer storage system must be sized according to the volume and quantity required for the process. Temperature control should be performed by CT. Weight measurement can be performed by a load cell or flow meter, and the system may include accessories such as an online pH meter, dissolved oxygen (DO) sensor, and conductivity sensor (CT). | O sistema de armazenamento de buffer deve ser dimensionado de acordo com o volume e a quantidade necessários para o processo. O controle de temperatura deve ser realizado por TC. A medição de peso pode ser realizada por uma célula de carga ou medidor de vazão, e o sistema pode incluir acessórios como um medidor de pH online, sensor de oxigênio dissolvido (OD) e sensor de condutividade (TC). | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
| IVT | IVT reactor |  |  | URS2, 5.1.81 | No content of parameter requirement indicated in URS | What is the Volume? | Between 10-50L | Entre 10-50L | Based on the normal practise, working volume of SUT fermenter is 30L. Then the assumption is Plasmid: 30L*0.6g/L=18g per batch Purified plasmid: 18g*0.4=7.2g per batch Linearization: 7.2g*0.85=6.12g per batch 6.12g*75 batch=459g per year ”Over 300 days, ~150 cycles can be run“: 459g/150=~3g “1.06 g of linear pDNA, yielding ~21.2 g of mRNA per cycle”: Plasmid convert to mRNA=1:20 So IVT batch size: 3g*20=60g mRNA
|  | Mixer |  |  |  | Intermediate material solution storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | Between 10-50L | Entre 10-50L | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | Bead- based purification system |  |  | URS2, 5.1.99 | No content of parameter requirement indicated in URS | The magnetic bead-based mRNA purification technology is suitable for small volume / laboratory-scale purification (1 - 1.5 mg of mRNA), and the chromatography method is recommended for large-scale production. Please confirm chromatography method is to be used in this project | Yes, you can opt for chromatographic purification. | Sim, você pode optar pela purificação cromatográfica.
|  | Chromatographic system |  |  | URS2, 5.1.100 | No content of parameter requirement indicated in URS | What is the Equipment quantity? ； Please check below requirement of the system: ☐Single - use chromatography system，☐Stainless steel chromatography system. ☐ Gradient function； ☐ Chromatography system: flow rate： L/h； Inlet quantity： Pcs， Outlet quantity： Pcs； ☐ Pre-filter； Detectors： ☐ Bubble sensor at the sample inlet；☐ Bubble sensor at the pre-column； ☐ UV sensor： pcs； Brand ; ☐ single wavelength, nm;☐ dual wavelength, nm; ☐ full wavelength, nm; ☐ Pressure sensor: pcs, location: ☐ post pump, ☐ pre-column, ☐ post-column; ☐ pH sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Conductivity sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Temperature sensor pcs, location . ☐ Explosion - proof function; The salt concentration of buffer and sample in each step mol，pH ； Other requirement: . | The chromatographic system should be specified according to the quantity of equipment required and must be single-use. It must have a gradient function and operate at a defined flow rate in L/h, with the number of inlets and outlets depending on the process. It is recommended to include a pre-filter and sensors such as: bubble sensors at the sample inlet and before the column, UV sensors (single-wavelength, fluorescence, dual-spectrum, or full-spectrum), pressure sensors (after the pump, before and after the column), pH and conductivity sensors (before and after the column), and temperature sensors in strategic locations. The system may include an explosion-proof function and must consider the salt concentration and pH of the buffers and samples at each stage. The brand and specifications of the sensors should be defined according to the process requirements. | O sistema cromatográfico deve ser especificado de acordo com a quantidade de equipamentos necessários e deve ser de uso único. Deve ter função de gradiente e operar a uma vazão definida em L/h, com o número de entradas e saídas dependendo do processo. Recomenda-se incluir um pré-filtro e sensores como: sensores de bolhas na entrada da amostra e antes da coluna, sensores UV (comprimento de onda único, fluorescência, espectro duplo ou espectro total), sensores de pressão (após a bomba, antes e depois da coluna), sensores de pH e condutividade (antes e depois da coluna) e sensores de temperatura em locais estratégicos. O sistema pode incluir função à prova de explosão e deve considerar a concentração de sal e o pH dos tampões e amostras em cada estágio. A marca e as especificações dos sensores devem ser definidas de acordo com os requisitos do processo.
|  | Column |  |  | URS2, 5.1.100 | No content of parameter requirement indicated in URS | Please check below requirement of the system: ☐Automatic；☐Manual； Column tube diameter mm； Column tube height mm； Column tube material: ☐Acrylic; ☐Stainless steel; Particle size of the chromatography resin μm； Mesh size: ☐10μm; ☐20μm; ☐Others: μm; Maximum operation pressure bar； Packing equipment： ☐ Automatic packing,☐ Manual packing； ☐ Automatic packing station; ☐ Slurry tank；volume: L; | You can consider the same as the previous one | Você pode considerar o mesmo que o anterior | Normally the resin will be used multiply. Customer accepts the column to br used multiply based on CIP. Customer also expects a quotation of resin price to evalute the operation cost if using the resn only one time for each batch.
|  | UF system | 0.5 to 2.5 m2 |  | URS2, 5.1.95 |  | Please check below requirement of the system: ☐Semi-automatic; ☐Automatic； Initial volume： L；Final volume L； Max. flow rate L/h； ☐Cassette，brand ，model ； ☐Hollow fiber，brand ，model ； System: ☐ Integrity test port reserved; ☐ Online integrity test module； ☐ Pressure sensor； ☐ Temperature sensor pcs，location: ☐ feed, ☐ retentate, ☐ permeate; ☐ pH sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Conductivity sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ UV sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Electromagnetic flowmeter, location: ☐ retentate, ☐ permeate; ☐ Mass flowmeter, location: ☐ retentate, ☐ permeate; ☐ Pipeline sampling valve, location: ☐ retentate, ☐ permeate; The salt concentration of the sample mol，pH ； ☐Organic reagents, concentration ； ☐Explosion - proof function； ☐TFF system circulating tank: Tank volume L，☐Weighing，☐Stirring，☐Tank thermal insulation layer；☐Temperature sensor; ☐Conductivity sensor; ☐ Sampling valve; ☐ Replenishment module: ☐ Replenishment inlet quantity ; Replenishment flow rate L/h； | The ultrafiltration (UF) system for concentration and diafiltration or microfiltration should preferably be automatic, with adjustable initial and final volumes depending on the process, and a maximum flow rate compatible with the production scale. It can use cassettes or hollow fibers, with an adequate membrane area (up to several m²), and include pressure, temperature, pH, conductivity, and UV sensors, as well as electromagnetic or mass flow meters, sampling valves, and an integrity testing module. The system can operate with organic reagents, buffers with defined saline concentrations and pH, and include a recirculating tank with agitation, thermal insulation, sensors, and a replenishment module with flow control. | O sistema de ultrafiltração (UF) para concentração e diafiltração ou microfiltração deve ser preferencialmente automático, com volumes inicial e final ajustáveis ​​dependendo do processo e vazão máxima compatível com a escala de produção. Pode utilizar cassetes ou fibras ocas, com área de membrana adequada (até vários m²), e incluir sensores de pressão, temperatura, pH, condutividade e UV, além de medidores de vazão eletromagnéticos ou mássicos, válvulas de amostragem e um módulo de teste de integridade. O sistema pode operar com reagentes orgânicos, tampões com concentrações salinas e pH definidos, e incluir um tanque de recirculação com agitação, isolamento térmico, sensores e um módulo de reabastecimento com controle de vazão.
|  | Mixer |  |  | URS2, 5.1.80 | No content of parameter requirement indicated in URS | What are the Volume and quantity? | Preferably dispensable step | Preferencialmente dispensavel | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | Sterilizing filtration system |  |  | URS2, 5.1.106 | No filter size indicated. | 1.The volume of the material to be processed? 2.Filtration flowrate? | Volume to be processed 0.5-10L at filtration rate 50L/h (maximum) can be optimized. | O volume a ser processado de 0,5-10 L a uma taxa de filtração de 50 L/h (máximo) pode ser otimizado.
|  | Storage system |  |  | URS2, 5.1.93 | No content of parameter requirement indicated in URS | What are the Volume and quantity? | 0,5-10L | 0,5-10L | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
| Add-cap | Wave |  |  |  | Add-cap | What is the Volume? | Consider the same as IVT | Considere o mesmo que IVT
|  | Mixer |  |  |  | Intermediate material solution storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor |  |  | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | Chromatographic system |  |  |  | Remove the uncapped mRNA and enzyme | What is the Equipment quantity? ； Please check below requirement of the system: ☐Single - use chromatography system，☐Stainless steel chromatography system. ☐ Gradient function； ☐ Chromatography system: flow rate： L/h； Inlet quantity： Pcs， Outlet quantity： Pcs； ☐ Pre-filter； Detectors： ☐ Bubble sensor at the sample inlet；☐ Bubble sensor at the pre-column； ☐ UV sensor： pcs； Brand ; ☐ single wavelength, nm;☐ dual wavelength, nm; ☐ full wavelength, nm; ☐ Pressure sensor: pcs, location: ☐ post pump, ☐ pre-column, ☐ post-column; ☐ pH sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Conductivity sensor： pcs, location: ☐ pre-column, ☐ post-column; ☐ Temperature sensor pcs, location . ☐ Explosion - proof function; The salt concentration of buffer and sample in each step mol，pH ； Other requirement: .
|  | Column |  |  |  |  | Please check below requirement of the system: ☐Automatic；☐Manual； Column tube diameter mm； Column tube height mm； Column tube material: ☐Acrylic; ☐Stainless steel; Particle size of the chromatography resin μm； Mesh size: ☐10μm; ☐20μm; ☐Others: μm; Maximum operation pressure bar； Packing equipment： ☐ Automatic packing,☐ Manual packing； ☐ Automatic packing station; ☐ Slurry tank；volume: L; |  |  | Normally the resin will be used multiply. Customer accepts the column to br used multiply based on CIP. Customer also expects a quotation of resin price to evalute the operation cost if using the resn only one time for each batch.
|  | UF system |  |  |  | Concentration and diafiltration | Please check below requirement of the system: ☐Semi-automatic; ☐Automatic； Initial volume： L；Final volume L； Max. flow rate L/h； ☐Cassette，brand ，model ； ☐Hollow fiber，brand ，model ； Max membrane area： m2; System: ☐ Integrity test port reserved; ☐ Online integrity test module； ☐ Pressure sensor； ☐ Temperature sensor pcs，location: ☐ feed, ☐ retentate, ☐ permeate; ☐ pH sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Conductivity sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ UV sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Electromagnetic flowmeter, location: ☐ retentate, ☐ permeate; ☐ Mass flowmeter, location: ☐ retentate, ☐ permeate; ☐ Pipeline sampling valve, location: ☐ retentate, ☐ permeate; The salt concentration of the sample mol，pH ； ☐Organic reagents, concentration ； ☐Explosion - proof function； ☐TFF system circulating tank: Tank volume L，☐Weighing，☐Stirring，☐Tank thermal insulation layer；☐Temperature sensor; ☐Conductivity sensor; ☐ Sampling valve; ☐ Replenishment module: ☐ Replenishment inlet quantity ; Replenishment flow rate L/h；
|  | Storage system |  |  |  | Buffer storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor |  |  | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
LNP | Buffer preparation | Preparation system | between 100L and 500L |  | URS3, 5.1.80-5.1.84 |  | 1.Whar are the Volume and quantity? 2.This step will require some Ethanol lipid solvent. The preparation and storage of this solvent need to consider explosion-proof design, and single-use system cannot be used. Please confirm Stainless Steel is acceptable. Please provide the quantity and volume of the specially designed preparation and storage tanks. | Between 100 and 500L, single-use only, and all materials in contact with the product must comply with this requirement. Ethanol is required and must be explosion-proof. The volume must be at least four times the volume to be encapsulated. | Entre 100 e 500 L, o produto é de uso único e todos os materiais em contato com o produto devem atender a este requisito. O etanol é obrigatório e deve ser à prova de explosão. O volume deve ser pelo menos quatro vezes maior que o volume a ser encapsulado. | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
| Entrapment | Microjet homogenizer | 48 L/h |  | URS3, 5.1.85-5.1.95 |  | Please provide below requirement of the system: 1. Chip flow rate; 2. Batch volumes of the aqueous phase solution and oil phase solution to be emulsified; 3.The ratio of the aqueous phase to the oil phase; 4.Time requirements for processing. | Flow rate required to obtain particles between 20 and 100 nm Volume to be defined by the manufacturer, meeting the demand of 1 L of mRNA solution for 4 liters of oil phase 1 liter of aqueous phase for 4 liters of oil phase Time between 1 and 6 hours to obtain particles between 20 and 100 nm | Vazão necessária para obtenção de partículas entre 20 e 100 nm Volume a ser definido pelo fabricante, atendendo à demanda de 1 L de solução de mRNA para 4 litros de fase oleosa 1 litro de fase aquosa para 4 litros de fase oleosa Tempo entre 1 e 6 horas para obtenção de partículas entre 20 e 100 nm
|  | Storage system |  |  |  | Buffer storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | Temperature control via TCU, no need to modify the atmosphere in this process, online pH/T°C meter, DO sensor, CT sensor depending on the process can vary between 100 - 500L | Controle de temperatura via TCU, sem necessidade de modificar a atmosfera neste processo, medidor de pH/T°C online, sensor de OD, sensor de CT dependendo do processo pode variar entre 100 - 500L | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
| Ultrafiltration | Mixer |  |  |  | Intermediate material solution storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | Same specifications as buffer mixer | Mesmas especificações do misturador de buffer | Customer agreed to provide the number of the mixers. | I believe that, 8 mixers can cover all steps , provided the schedule is optimized to reuse mixers for non-simultaneous steps. For greater robustness, especially with high batch frequency or simultaneous buffer needs, eight mixers are ideal, configured as: 2 mixers (1–10 L) for linearization, digestion, and intermediate storage. 3 mixers (10–50 L) for lysis, TFF, chromatography, media and buffer preparation, and IVT. 3 mixers (100–500 L) for LNP, larger buffers, and oil phase.
|  | UF system | filtering areas 1 to 10m2 |  | URS3, 5.1.96 |  | Please check below requirement of the system: ☐Semi-automatic; ☐Automatic； Initial volume： L；Final volume L； Max. flow rate L/h； ☐Cassette，brand ，model ； ☐Hollow fiber，brand ，model ； Max membrane area： m2; System: ☐ Integrity test port reserved; ☐ Online integrity test module； ☐ Pressure sensor； ☐ Temperature sensor pcs，location: ☐ feed, ☐ retentate, ☐ permeate; ☐ pH sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Conductivity sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ UV sensor： pcs，location: ☐ retentate, ☐ permeate; ☐ Electromagnetic flowmeter, location: ☐ retentate, ☐ permeate; ☐ Mass flowmeter, location: ☐ retentate, ☐ permeate; ☐ Pipeline sampling valve, location: ☐ retentate, ☐ permeate; The salt concentration of the sample mol，pH ； ☐Organic reagents, concentration ； ☐Explosion - proof function； ☐TFF system circulating tank: Tank volume L，☐Weighing，☐Stirring，☐Tank thermal insulation layer；☐Temperature sensor; ☐Conductivity sensor; ☐ Sampling valve; ☐ Replenishment module: ☐ Replenishment inlet quantity ; Replenishment flow rate L/h； | The ultrafiltration (UF) system for concentration and diafiltration or microfiltration should preferably be automatic, with adjustable initial and final volumes depending on the process, and a maximum flow rate compatible with the production scale. It can use cassettes or hollow fibers, with an adequate membrane area (up to several m²), and include pressure, temperature, pH, conductivity, and UV sensors, as well as electromagnetic or mass flow meters, sampling valves, and an integrity testing module. The system can operate with organic reagents, buffers with defined saline concentrations and pH, and include a recirculating tank with agitation, thermal insulation, sensors, and a replenishment module with flow control. | O sistema de ultrafiltração (UF) para concentração e diafiltração ou microfiltração deve ser preferencialmente automático, com volumes inicial e final ajustáveis ​​dependendo do processo e vazão máxima compatível com a escala de produção. Pode utilizar cassetes ou fibras ocas, com área de membrana adequada (até vários m²), e incluir sensores de pressão, temperatura, pH, condutividade e UV, além de medidores de vazão eletromagnéticos ou mássicos, válvulas de amostragem e um módulo de teste de integridade. O sistema pode operar com reagentes orgânicos, tampões com concentrações salinas e pH definidos, e incluir um tanque de recirculação com agitação, isolamento térmico, sensores e um módulo de reabastecimento com controle de vazão.
|  | Storage system |  |  |  | Buffer storage | 1.What are the Volume and quantity? Please check below requirement of the system: 2.Process air: ☐nitrogen ☐compressed air ☐vacuum 3.Temperature control medium：☐ Chilled water ☐cooling water ☐hot water ☐plant steam ☐temperature control unit 4.weight mode：☐load cell ☐flowmeter 5.accessories：☐online pH meter ☐ DO sensor ☐CT sensor | Same specification as the LNP buffer, maintaining biocompatibility characteristics for use (pH, viscosity and osmolarity) | Mesma especificação do tampão LNP, mantendo as características de biocompatibilidade para uso (pH, viscosidade e osmolaridade) | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
| Sterilization and filtration | Sterilizing filtration system |  |  | URS3, 5.1.101-5.1.104 | No filter size indicated. | 1.What is the volume of the material to be processed? 2.Filtration flowrate? | 0.5-50L and filtration rate up to 50L/h | 0,5-50L e taxa de filtração de até 50L/h
|  | Storage system |  |  |  |  | What is the Volume? | 0.5-50L depending on the vaccine or demand | 0,5-50L dependendo da vacina ou demanda | Customer agreed to provide the number of the storage tank. | Analysis by Stage pDNA Unit Operations: Fermentation (50 L): Product (~50 L, 1.5 g/L dry biomass) temporarily stored post-fermentation. Alkaline Lysis: ~50 L of lysed material (1.5 g/L dry biomass). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume post-concentration (~10–50 L). Retentate and permeate may require temporary storage. Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~10–50 L) and waste. Linearization: ~5 L (1/10 of fermentation volume). Buffers: TFF: 1 buffer (~10–50 L) for diafiltration/washing. Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 total buffers (~10–50 L each, depending on column volume and flow). Required Tanks: 1 tank (50 L) for post-fermentation. 1 tank (50 L) for post-lysis. 1 tank (10–50 L) for post-TFF (concentrate). 1 tank (5–10 L) for post-linearization. 2 tanks (10–50 L) for chromatography fractions (one per type, assuming sequential collection). 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~14 tanks (1x 5–10 L, 11x 10–50 L, 2x 50 L). IVT Unit Operations: IVT Reaction (50 L, working volume ~30 L): Product (~30 L, ~21.2 g mRNA/cycle). TFF (Microfiltration/Ultrafiltration/Concentration): Reduced volume (~0.5–10 L post-concentration). Chromatography (Ion Exchange and Hydrophobic Interaction): Collected fractions (~0.5–10 L). Buffers: TFF: 1 buffer (~10–50 L). Chromatography: 4 buffers per chromatography (ion exchange + hydrophobic interaction) = 8 buffers (~10–50 L each). Required Tanks: 1 tank (30–50 L) for post-IVT. 1 tank (0.5–10 L) for post-TFF. 2 tanks (0.5–10 L) for chromatography fractions. 8 tanks (10–50 L) for chromatography buffers. 1 tank (10–50 L) for TFF buffer. Total: ~13 tanks (3x 0.5–10 L, 9x 10–50 L, 1x 30–50 L). LNP Unit Operations: Microfluidization: Aqueous phase (~1 L mRNA) + oil phase (~4 L, total volume up to 500 L to support 1:4 ratio). Post-Formulation Concentration (TFF): Final volume ~0.5–50 L. Buffers: TFF: 1 buffer (~100–500 L, due to larger oil phase volume). Other Buffers: 1–2 additional buffers (~100–500 L) for phase preparation and post-formulation. Required Tanks: 1 tank (100–500 L) for aqueous/oil phase. 1 tank (0.5–50 L) for post-TFF. 1 tank (100–500 L) for TFF buffer. 2 tanks (100–500 L) for additional buffers. Total: ~5 tanks (1x 0.5–50 L, 4x 100–500 L).
Customer agreed to send the facility layout with indication of area to be used for fermentation, IVT, LNP respectively